## Protocol Summary
Summary of (proposed) c-VEDA Imaging Protocol
| Breakdown tab start: | 12 |  |  |  |  | 16 | 12 | 13 | 14 | 15 | 21 | 38
| Param tab start: | 11 |  |  |  |  | 28 | 24 | 25 | 26 | 27 | 33 | 50
Acquired
|  |  |  | Sequence | Scan duration (approx) | Running session time (inc current scan) | 2D/3D | Voxel size: freq dirn (mm) | Voxel size: phase dirn (mm) | Voxel size: slice dirn (mm) | Slice gap (mm) | Slice acquisition direction | PI reduction factor
Structural
|  | row number | Setup
|  |  |  | Participant positioning, etc | 0:10 | 0:10
|  |  |  | Localiser and slice prescription | 0:01 | 0:11
|  |  |  | Parallel imaging calibration | 0:00 | 0:12
|  |  |  | Pre-scans (centre frequency, shimming etc) | 0:02 | 0:14
| 1 | 11 |  | 2D T2-weighted FSE/TSE | 0:01 | 0:15 | 2D | 0.75 | 0.75 | 4 | 0 | I to S | 2
| 2 | 71 |  | 2D (fast) FLAIR | 0:02 | 0:18 | 2D | 0.9375 | 1.875 | 4 | 0 | I to S | 0
| 3 | 131 |  | 3D T1-weighted volume | 0:10 | 0:27 | 3D | 1 | 1 | 1.2 | 0 | L to R | 0
| 4 | 191 |  | DTI | 0:09 | 0:36 | 2D | 2.4 | 2.4 | 2.4 | 0 | I to S | 2
b = 1300, 32 directions b = 0, 4 measurements
| 5 | 251 |  | DTI (reversed) | 0:02 | 0:39 | 2D | 2.4 | 2.4 | 2.4 | 0 | I to S | 2
b = 1300, 6 directions b = 0, 4 measurements
Functional
Setup
|  |  |  | Explanation of paradigms to participant | 0:02 | 0:41
| 6 | 311 |  | Resting State fMRI | 0:06 | 0:47 | 2D | 3.4 | 3.4 | 2.4 | 1 | S to I | 2
164 volumes
Other
| 7 | 371 |  | B0 mapping | 0:01 | 0:48 | 0:00 | 4 | 4 | 4 | 0 | 0:00 | 0
|  |  |  | Total estimated protocol duration | 0:48

## Seq, Param & Manu Desc
Sequence List
|  | Imaging Type | Sequence | GE | Siemens | Philips
|  | T1 Weighted | Inversion Prepared Fast Spoiled Gradient Echo (3D) | IR-FSPGR | MPRAGE | Inversion prepared TFE
|  | DTI | Dual Spin Echo EPI | Dual Spin Echo EPI | Twice Refocused EPI | -
|  | T2 Weighted | Fast/Turbo Spin Echo (2D) | FSE | TSE | TSE
|  | fMRI | Gradient Echo EPI | GRE-EPI | GE-EPI | FFE-EPI
Sequence and Parameter Description, Manufacturer Terminology
Parameter List
| Category | Parameter | Description | GE | Siemens | Philips
0 | Timing | TR (ms) | Repetition Time
1 |  | TE (ms) | Echo Time
2 |  | TI (ms) | Inversion Time
3 |  | NSA | Number of signal averages (I.e. number of times same line in k-space sampled) | NEX | NSA | ACQ
4 |  | Scan duration | Scan time reported by scanner
5 |  | Peripheral gated (Y/N) | In this sequence, is peripheral gating performed?
6 | Resolution | Matrix size: in-plane, freq dirn | Number of points acquired in frequency encode direction
7 |  | Matrix size: in-plane, phase dirn | Number of points acquired in in-plane phase encode direction
8 |  | Total No. of slices (2D) or Matrix size: slice dirn (3D) | Number of slices acquired (2D) or number of phase encoding steps in slice direction (3D) | No. of slices(2D) Locs per slab(3D)
9 |  | FOV frequency (cm) | Field of view in frequency encode direction (cm) |  | FOV read
10 |  | FOV phase (%) | Field of view in phase encode direction as percentage of frequency FOV (keeping voxel size I.e. total kspace coverage constant) | Phase FOV (0 to 1) | FOV phase (%)
11 |  | FOV slice (cm) | Coverage in the slice/slab direction
12 |  | Voxel size: freq dirn (mm) | Voxel size in frequency encode direction (mm) |  | Base resolution
13 |  | Voxel size: phase dirn (mm) | Voxel size in phase encode direction (mm) |  | Phase resolution
14 |  | Voxel size: slice dirn (mm) | Slice thickness (2D) or slice direction voxel size (3D) | Slice thickness | Slice resolution
15 |  | Slice gap (mm) | Gap between slices (2D acquisition) | spacing
16 | Coverage | 2D/3D | 2D or 3D acquisition mode
17 |  | In-plane phase enc. dirn. (AP,LR,SI) | Direction of phase encoding within slice
18 |  | Region of slice coverage | Anatomical region scanned - always whole head
19 |  | Slice orientation (Sag/Ax/Cor/Obl) | Acquisition of slice orientation (Sagittal/Axial/Coronal/Oblique - e.g. ACPC)
20 |  | Slice order (Inter/Seq) | Order of slice acquisition (Interleaved/Sequential) |  | Slice mode
21 |  | Slice acquisition direction | Direction of slice acquisition e.g. Sup>Inf, Post>Ant etc |  | Series
22 |  | K-space ordering | Order in which k-space lines are collected | View ordering ("top down" versus "bottom up") | Phase encode direction (A>>P versus P>>A) | Fat Shift Direction ("A" versus "P")
23 |  | Concatenations | Use of several blocks of interleaved slices | multiple acquisitions | Concatenations | multiple packages
24 |  | Oversampling Phase (%) | Increase in FOV in the in-plane phase encode direction, to place aliased signal outside the actual desired FOV | No phase wrap
25 |  | Oversampling Slice dirn 3D (%) | Increase in FOV in the slice phase encode direction (3D), to place aliased signal outside the actual desired FOV
26 |  | Slice overlap (mm) | Slice overlap (2D acquisition interleaved) e.g. Slice thickness of 4 mm acquired but with 3mm between centre of slices = slice overlap 1mm | over contiguous
27 | Sequence specific | ETL/ E Spacing (us) | Echo train length (Turbo/Fast spin echo sequence). Echo Spacing - time between each echo in the train |  | Turbo factor (Combination of ETL AND Partial Fourier factor)
28 |  | TSE corrections | Echo train is reversed and averaged | No Equivalent - should be switched off for other manufacturers | Compensate T2 decay
29 |  | Fat Sat (CHESS or Equiv) on (Y/N) | Frequency selective fat saturation prepared sequence |  |  | SPIR
30 | Other acquisition | Partial Fourier (Y/N) | Partial k-space coverage
31 |  | Partial Fourier method (NEX/Echo) | Fractional NEX , Fractional Echo (see description below)
32 |  | Fractional NEX | Full kx, fractional ky (I.e. full echoes collected with less lines) | NEX
33 |  | Fractional Echo | Fractional kx, Full ky (I.e. collect partial echoes of all lines) | "Minimum TE"
34 |  | Flip angle (degrees) | RF Flip angle
35 |  | rBW (Hz per pixel) | Receiver bandwidth (per pixel; total is this value * (Matrix size: in-plane, freq dirn))
36 |  | rBW (kHz) | Receiver bandwidth (total bandwidth is +/- this value)
37 | Parallel imaging | PI method | Parallel imaging acquisition method | ASSET | iPAT, GRAPPA (kspace) mSENSE (image space) | SENSE
38 |  | PI reduction factor | Parallel imaging reduction factor | ASSET Factor |  | SENSE factor
39 | Corrections | B0 mapping (Y/N) | Mapping of B0 inhomogeneity for this sequence?
40 |  | B1 mapping: Transmit (Y/N) | Mapping of B1 Transmit inhomogeneity for this sequence?
41 |  | B1 mapping: Receive (Y/N) | Mapping of B1 Receive inhomogeneity for this sequence?
42 |  | Grad Non-lin correction (Y/N) | Apply on-line gradient non-linearity correction to data in this sequence?
43 |  | Grad Non-lin correction Method | Method of Gradient non-linearity correction | GRADwarp | Large FOV filter
44 | Pre-processing | Zero filling (Y/N) | Acquired matrix is zero-filled to reconstruct to a larger matrix size
45 |  | Final Image matrix size | Size of final image, if different from the acquired matrix
46 |  | Filtering | A range of paramters affecting the "smoothness" of the final image | changed by setting CVs rhfermr; rhfermw | Values are: "weak", "medium" or "strong" | no easily adjustable
47 | Diffusion specific | b value (s mm-2) | b value of diffusion weighting
48 |  | No. dirns | Number of diffusion weighted directions
49 |  | No. of b=0 acquisitions | Number of b = 0 (or T2) acquisitions
50 |  | Direction scheme (1/2/3) | Orientation of directions (1 = default supplied by manufacturer, 2 = Own bespoke, 3 = Published/Other recognised scheme e.g distributed charges on sphere [D Jones Magn Reson Med 1999])
51 |  | Eddy current reduction technique (Y/N) | Eddy current artefacts nulled using alternative diffusion gradient arrangement and double spin echo. [O. Heid ISMRM 2000, TG. Reese Magn Reson Med 2003] | Dual Spin Echo | Twice refocusing | -
52 | fMRI specific | No. vols. | Number of samples/volumes (typically 100-200) collected for an fMRI paradigm | Samples | Measurements | Dyn Scans
53 |  | No. dummy scans | Number of samples/volumes (typically 3-4) run without data collection at the start of an fMRI run, to allow longitudinal magnetisation to reach a steady state | Dummy samples |  | Dummy Scans
| Other | Orientations | Terms used to describe orientations of slices:
|  |  | Superior | Superior | Head | Head
|  |  | Inferior | Inferior | Feet | Feet
|  |  | Anteroir | Anteroir | Anteroir | Anteroir
|  |  | Posterior | Posterior | Posterior | Posterior
|  |  | Left | Left | Left | Left
|  |  | Right | Right | Right | Right
|  |  | Axial | Axial | Transversal | Transverse
|  |  | Coronal | Coronal | Coronal | Coronal
|  |  | Sagittal | Sagittal | Sagittal | Sagittal

## Sequence Breakdown - To fill in
Sequence Break Down by Parameter
Structural
|  |  | DTI |  |  |  |  | Site Name:
fMRI
Other
Y | [6 empty] | Can you match "ideal"? (Y/N) | If answered No, what is the next best value you can achieve? If this "best" parameter is influenced by another please describe in comments section. | Comments
N
|  | Imaging Type | Sequence | Parameter Category | Sequence Parameter | "Ideal"
| 1 | Structural | 2D T2-weighted FSE/TSE | Timing | TR (ms) | 4380 |  | 4920
[5 empty] | TE (ms) | 60 |  | 69
TI (ms)
[5 empty] | NSA | 1 |  | 1
[5 empty] | Scan duration | 0:01 |  | 1:02
[5 empty] | Peripheral gated (Y/N) | N |  | N
|  |  |  | Resolution | Matrix size: in-plane, freq dirn | 320 |  | 320
[5 empty] | Matrix size: in-plane, phase dirn | 320 |  | 320
[5 empty] | Total No. of slices (2D) or Matrix size: slice dirn (3D) | 36 |  | 36
[5 empty] | FOV frequency (cm) | 24 |  | 24
[5 empty] | FOV phase (%) | 75 |  | 75
[5 empty] | FOV slice (cm) | 14.4
[5 empty] | Voxel size: freq dirn (mm) | 0.75 |  | 0.8
[5 empty] | Voxel size: phase dirn (mm) | 0.75 |  | 0.8
[5 empty] | Voxel size: slice dirn (mm) | 4 |  | 4
[5 empty] | Slice gap (mm) | 0
|  |  |  | Coverage | 2D/3D | 2D |  | 2D
[5 empty] | In-plane phase enc. dirn. (AP,LR,SI) | LR |  | RL
[5 empty] | Region of slice coverage | Whole Brain |  | WHOLE BRAIN
[5 empty] | Slice orientation (Sag/Ax/Cor/Obl) | Oblique AC-PC |  | OBLIQUE AC PC
[5 empty] | Slice order (Inter/Seq) | Interleaved |  | INTERLEAVED
[5 empty] | Slice acquisition direction | I to S |  | I TO S
K-space ordering
[5 empty] | Concatenations | 2 |  | 2
Oversampling Phase (%)
Oversampling Slice dirn 3D (%)
Slice overlap (mm)
|  |  |  | Sequence specific | ETL/ E Spacing (us) | 19 |  | 19
[5 empty] | TSE corrections | None |  |  | could not find this option
[5 empty] | Fat Sat (CHESS or Equiv) on (Y/N) | N |  | N
|  |  |  | Other acquisition | Partial Fourier (Y/N) | N |  | N
Partial Fourier method (NEX/Echo)
Fractional NEX
Fractional Echo
[5 empty] | Flip angle (degrees) | 90 |  | 90
[5 empty] | rBW (Hz per pixel) | 156.25 |  | 156
[5 empty] | rBW (kHz) | 25
|  |  |  | Parallel imaging | PI method |  |  | GRAPPA
[5 empty] | PI reduction factor | 2 |  | 2
|  |  |  | Corrections | B0 mapping (Y/N)
B1 mapping: Transmit (Y/N)
B1 mapping: Receive (Y/N)
Grad Non-lin correction (Y/N)
Grad Non-lin correction Method
|  |  |  | Pre-processing | Zero filling (Y/N) | Y |  |  | could not find this option
[5 empty] | Final Image matrix size | 512
Filtering
|  |  |  | Diffusion specific | b value (s mm-2)
No. dirns
No. of b=0 acquisitions
Direction scheme (1/2/3)
Eddy current reduction technique (Y/N)
|  |  |  | fMRI specific | No. vols.
No. dummy scans
| 2 | Structural | 2D (fast) FLAIR | Timing | TR (ms) | 8000 |  | 8000
[5 empty] | TE (ms) | 120 |  | 117
[5 empty] | TI (ms) | 2000 |  | 2000
[5 empty] | NSA | 1 |  | 1
[5 empty] | Scan duration | 0:02 |  | 2.25
[5 empty] | Peripheral gated (Y/N) | N |  | N
|  |  |  | Resolution | Matrix size: in-plane, freq dirn | 256 |  | 256
[5 empty] | Matrix size: in-plane, phase dirn | 128 |  | 50%
[5 empty] | Total No. of slices (2D) or Matrix size: slice dirn (3D) | 36 |  | 36
[5 empty] | FOV frequency (cm) | 24 |  | 24
[5 empty] | FOV phase (%) | 100 |  | 100
[5 empty] | FOV slice (cm) | 14.4
[5 empty] | Voxel size: freq dirn (mm) | 0.9375 |  | 0.9
[5 empty] | Voxel size: phase dirn (mm) | 1.875 |  | 0.9
[5 empty] | Voxel size: slice dirn (mm) | 4 |  | 4
[5 empty] | Slice gap (mm) | 0
|  |  |  | Coverage | 2D/3D | 2D |  | 2D
[5 empty] | In-plane phase enc. dirn. (AP,LR,SI) | LR |  | RL
[5 empty] | Region of slice coverage | Whole Brain |  | WHOLE BRAIN
[5 empty] | Slice orientation (Sag/Ax/Cor/Obl) | Oblique AC-PC |  | OBLIQUE AC PC
[5 empty] | Slice order (Inter/Seq) | Interleaved |  | INTERLEAVED
[5 empty] | Slice acquisition direction | I to S |  | I TO S
K-space ordering
[5 empty] | Concatenations | 3 |  | 2
Oversampling Phase (%)
Oversampling Slice dirn 3D (%)
[5 empty] | Slice overlap (mm) | 0 |  | 0
|  |  |  | Sequence specific | ETL/ E Spacing (us)
[5 empty] | TSE corrections | N |  | N
[5 empty] | Fat Sat (CHESS or Equiv) on (Y/N) | N |  | N
|  |  |  | Other acquisition | Partial Fourier (Y/N) | N |  | N
Partial Fourier method (NEX/Echo)
Fractional NEX
Fractional Echo
[5 empty] | Flip angle (degrees) | 180 / 90 |  | 180
[5 empty] | rBW (Hz per pixel) | 195.3125 |  | 195
[5 empty] | rBW (kHz) | 25
|  |  |  | Parallel imaging | PI method | None |  | NONE
PI reduction factor
|  |  |  | Corrections | B0 mapping (Y/N)
B1 mapping: Transmit (Y/N)
B1 mapping: Receive (Y/N)
Grad Non-lin correction (Y/N)
Grad Non-lin correction Method
|  |  |  | Pre-processing | Zero filling (Y/N) | Y
[5 empty] | Final Image matrix size | 256
Filtering
|  |  |  | Diffusion specific | b value (s mm-2)
No. dirns
No. of b=0 acquisitions
Direction scheme (1/2/3)
Eddy current reduction technique (Y/N)
|  |  |  | fMRI specific | No. vols.
No. dummy scans
| 3 | Structural | 3D T1-weighted volume | Timing | TR (ms) |  |  | 1900
[5 empty] | TE (ms) |  |  | 2.43
[5 empty] | TI (ms) |  |  | 900
[5 empty] | NSA | 1 |  | 1
[5 empty] | Scan duration | 0:10
[5 empty] | Peripheral gated (Y/N) | N
|  |  |  | Resolution | Matrix size: in-plane, freq dirn | 256 |  | 256
[5 empty] | Matrix size: in-plane, phase dirn | 256 |  | 256
Total No. of slices (2D) or Matrix size: slice dirn (3D)
[5 empty] | FOV frequency (cm) | 26 |  | 26
[5 empty] | FOV phase (%) | 1 |  | 1
FOV slice (cm)
[5 empty] | Voxel size: freq dirn (mm) | 1 |  | 1
[5 empty] | Voxel size: phase dirn (mm) | 1 |  | 1
[5 empty] | Voxel size: slice dirn (mm) | 1.2 |  | 1
Slice gap (mm)
|  |  |  | Coverage | 2D/3D | 3D |  | 3D
[5 empty] | In-plane phase enc. dirn. (AP,LR,SI) | AP |  | AP
[5 empty] | Region of slice coverage | Whole Head |  | WHOLE HEAD
[5 empty] | Slice orientation (Sag/Ax/Cor/Obl) | Sagittal |  | SAGITTAL
Slice order (Inter/Seq)
[5 empty] | Slice acquisition direction | L to R |  | L TO R
K-space ordering
Concatenations
Oversampling Phase (%)
Oversampling Slice dirn 3D (%)
Slice overlap (mm)
|  |  |  | Sequence specific | ETL/ E Spacing (us)
TSE corrections
[5 empty] | Fat Sat (CHESS or Equiv) on (Y/N) | N |  | N
|  |  |  | Other acquisition | Partial Fourier (Y/N) | N |  | N
Partial Fourier method (NEX/Echo)
Fractional NEX
Fractional Echo
Flip angle (degrees)
rBW (Hz per pixel)
rBW (kHz)
|  |  |  | Parallel imaging | PI method | None |  | NONE
PI reduction factor
|  |  |  | Corrections | B0 mapping (Y/N)
B1 mapping: Transmit (Y/N)
B1 mapping: Receive (Y/N)
Grad Non-lin correction (Y/N)
Grad Non-lin correction Method
|  |  |  | Pre-processing | Zero filling (Y/N) | N
Final Image matrix size
Filtering
|  |  |  | Diffusion specific | b value (s mm-2)
No. dirns
No. of b=0 acquisitions
Direction scheme (1/2/3)
Eddy current reduction technique (Y/N)
|  |  |  | fMRI specific | No. vols.
No. dummy scans
| 4 | DTI | DTI | Timing | TR (ms) | 10000 - 150000 |  | 15000
[5 empty] | TE (ms) | 70-110 |  | 95
TI (ms)
[5 empty] | NSA | 1 |  | 1
[5 empty] | Scan duration | 0:09 |  | 9:32
[5 empty] | Peripheral gated (Y/N) | Y |  | N
|  |  |  | Resolution | Matrix size: in-plane, freq dirn | 128 |  | 128
[5 empty] | Matrix size: in-plane, phase dirn | 128 |  | 128
[5 empty] | Total No. of slices (2D) or Matrix size: slice dirn (3D) | 60 |  | 60
[5 empty] | FOV frequency (cm) | 30.7 |  | 30.7
[5 empty] | FOV phase (%) | 100 |  | 100
FOV slice (cm)
[5 empty] | Voxel size: freq dirn (mm) | 2.4 |  | 2.4
[5 empty] | Voxel size: phase dirn (mm) | 2.4 |  | 2.4
[5 empty] | Voxel size: slice dirn (mm) | 2.4 |  | 2.4
[5 empty] | Slice gap (mm) | 0 |  | 0
|  |  |  | Coverage | 2D/3D | 2D |  | 2D
[5 empty] | In-plane phase enc. dirn. (AP,LR,SI) | AP |  | AP
[5 empty] | Region of slice coverage | Whole Head |  | WHOLE HEAD
[5 empty] | Slice orientation (Sag/Ax/Cor/Obl) | Axial |  | AXIAL
[5 empty] | Slice order (Inter/Seq) | Interleaved |  | INTERLEAVED
[5 empty] | Slice acquisition direction | I to S |  | I TO S
[5 empty] | K-space ordering | P>>A |  | P>>A
Concatenations
Oversampling Phase (%)
Oversampling Slice dirn 3D (%)
Slice overlap (mm)
|  |  |  | Sequence specific | ETL/ E Spacing (us) | 48/0.612 |  | 0.72
TSE corrections
Fat Sat (CHESS or Equiv) on (Y/N)
|  |  |  | Other acquisition | Partial Fourier (Y/N) | Y |  | Y
[5 empty] | Partial Fourier method (NEX/Echo) | NEX |  | PHASE
[5 empty] | Fractional NEX | 0.75 |  | 0.75
Fractional Echo
Flip angle (degrees)
rBW (Hz per pixel)
rBW (kHz)
|  |  |  | Parallel imaging | PI method
[5 empty] | PI reduction factor | 2 |  | 2
|  |  |  | Corrections | B0 mapping (Y/N)
B1 mapping: Transmit (Y/N)
B1 mapping: Receive (Y/N)
[5 empty] | Grad Non-lin correction (Y/N) | Y |  | DYNAMIC FIELD CORRECTION
Grad Non-lin correction Method
|  |  |  | Pre-processing | Zero filling (Y/N) | N
Final Image matrix size
[5 empty] | Filtering | Strong |  | ELLIPTICAL FILTER
|  |  |  | Diffusion specific | b value (s mm-2) | 1300 |  | 1300
[5 empty] | No. dirns | 32 |  | 30
[5 empty] | No. of b=0 acquisitions | 3 |  | 3
[5 empty] | Direction scheme (1/2/3) | 1
[5 empty] | Eddy current reduction technique (Y/N) | Y |  | DYNAMIC FIELD CORRECTION
|  |  |  | fMRI specific | No. vols.
No. dummy scans
| 5 | DTI | DTI (reversed) | Timing | TR (ms) | 10000 - 150000 |  | 15000
[5 empty] | TE (ms) | 70-110 |  | 95
TI (ms)
[5 empty] | NSA | 1 |  | 1
[5 empty] | Scan duration | 0:02 |  | 0:04
[5 empty] | Peripheral gated (Y/N) | Y |  | N
|  |  |  | Resolution | Matrix size: in-plane, freq dirn | 128 |  | 128
[5 empty] | Matrix size: in-plane, phase dirn | 128 |  | 128
[5 empty] | Total No. of slices (2D) or Matrix size: slice dirn (3D) | 60 |  | 60
[5 empty] | FOV frequency (cm) | 30.7 |  | 30.7
[5 empty] | FOV phase (%) | 100 |  | 100
FOV slice (cm)
[5 empty] | Voxel size: freq dirn (mm) | 2.4 |  | 2.4
[5 empty] | Voxel size: phase dirn (mm) | 2.4 |  | 2.4
[5 empty] | Voxel size: slice dirn (mm) | 2.4 |  | 2.4
[5 empty] | Slice gap (mm) | 0 |  | 0
|  |  |  | Coverage | 2D/3D | 2D |  | 2D
[5 empty] | In-plane phase enc. dirn. (AP,LR,SI) | AP |  | AP
[5 empty] | Region of slice coverage | Whole Head |  | WHOLE HEAD
[5 empty] | Slice orientation (Sag/Ax/Cor/Obl) | Axial |  | AXIAL
[5 empty] | Slice order (Inter/Seq) | Interleaved |  | INTERLEAVED
[5 empty] | Slice acquisition direction | I to S |  | I TO S
[5 empty] | K-space ordering | A>>P |  | A>>P
Concatenations
Oversampling Phase (%)
Oversampling Slice dirn 3D (%)
Slice overlap (mm)
|  |  |  | Sequence specific | ETL/ E Spacing (us) | 48/0.612 |  | 0.72
TSE corrections
Fat Sat (CHESS or Equiv) on (Y/N)
|  |  |  | Other acquisition | Partial Fourier (Y/N) | Y |  | Y
[5 empty] | Partial Fourier method (NEX/Echo) | NEX |  | PHASE
[5 empty] | Fractional NEX | 0.75 |  | 0.75
Fractional Echo
Flip angle (degrees)
rBW (Hz per pixel)
rBW (kHz)
|  |  |  | Parallel imaging | PI method
[5 empty] | PI reduction factor | 2 |  | 2
|  |  |  | Corrections | B0 mapping (Y/N)
B1 mapping: Transmit (Y/N)
B1 mapping: Receive (Y/N)
[5 empty] | Grad Non-lin correction (Y/N) | Y |  | DYNAMIC FIELD CORRECTION
Grad Non-lin correction Method
|  |  |  | Pre-processing | Zero filling (Y/N) | N
Final Image matrix size
[5 empty] | Filtering | Strong |  | ELLIPTICAL FILTER
|  |  |  | Diffusion specific | b value (s mm-2) | 1300 |  | 1300
[5 empty] | No. dirns | 6 |  | 6
[5 empty] | No. of b=0 acquisitions | 3 |  | 3
[5 empty] | Direction scheme (1/2/3) | 1
[5 empty] | Eddy current reduction technique (Y/N) | Y |  | DYNAMIC FIELD CORRECTION
0
0
| 6 | fMRI | Resting State fMRI | Timing | TR (ms) | 2200 |  | 2200
[5 empty] | TE (ms) | 30 |  | 20
TI (ms)
[5 empty] | NSA | 1 |  | 1
[5 empty] | Scan duration | 0:06 |  | 0:06
[5 empty] | Peripheral gated (Y/N) | N |  | N
|  |  |  | Resolution | Matrix size: in-plane, freq dirn | 64 |  | 64
[5 empty] | Matrix size: in-plane, phase dirn | 64 |  | 64
[5 empty] | Total No. of slices (2D) or Matrix size: slice dirn (3D) | 40 |  | 40
[5 empty] | FOV frequency (cm) | 21.8 |  | 21.8
[5 empty] | FOV phase (%) | 100 |  | 100
[5 empty] | FOV slice (cm) | 13.2
[5 empty] | Voxel size: freq dirn (mm) | 3.4 |  | 3.4
[5 empty] | Voxel size: phase dirn (mm) | 3.4 |  | 3.4
[5 empty] | Voxel size: slice dirn (mm) | 2.4 |  | 2.4
[5 empty] | Slice gap (mm) | 1 |  | 1.2
|  |  |  | Coverage | 2D/3D | 2D |  | 2D
[5 empty] | In-plane phase enc. dirn. (AP,LR,SI) | AP |  | AP
[5 empty] | Region of slice coverage | Whole Brain |  | WHOLE BRAIN
[5 empty] | Slice orientation (Sag/Ax/Cor/Obl) | Oblique Axial: AC-PC |  | OBLIQUE AC PC
[5 empty] | Slice order (Inter/Seq) | Sequential |  | Descending
[5 empty] | Slice acquisition direction | S to I |  | S TO I
[5 empty] | K-space ordering | P>>A |  | P>>A
[5 empty] | Concatenations | 1 |  | 1
Oversampling Phase (%)
Oversampling Slice dirn 3D (%)
Slice overlap (mm)
|  |  |  | Sequence specific | ETL/ E Spacing (us) | 32/ |  | 0.65
TSE corrections
[5 empty] | Fat Sat (CHESS or Equiv) on (Y/N) | Y |  | Y
|  |  |  | Other acquisition | Partial Fourier (Y/N) | N |  | N
Partial Fourier method (NEX/Echo)
Fractional NEX
Fractional Echo
[5 empty] | Flip angle (degrees) | 75 |  | 75
[5 empty] | rBW (Hz per pixel) | 2000 |  | 2004
[5 empty] | rBW (kHz) | 64
|  |  |  | Parallel imaging | PI method |  |  | GRAPPA
[5 empty] | PI reduction factor | 2 |  | 2
|  |  |  | Corrections | B0 mapping (Y/N)
B1 mapping: Transmit (Y/N)
B1 mapping: Receive (Y/N)
Grad Non-lin correction (Y/N)
Grad Non-lin correction Method
|  |  |  | Pre-processing | Zero filling (Y/N) | N
Final Image matrix size
[5 empty] | Filtering | Weak
|  |  |  | Diffusion specific | b value (s mm-2)
No. dirns
No. of b=0 acquisitions
Direction scheme (1/2/3)
Eddy current reduction technique (Y/N)
|  |  |  | fMRI specific | No. vols. | 164 |  | 164
[5 empty] | No. dummy scans | 4 |  | 4
| 7 | Other | B0 mapping | Timing | TR (ms) |  |  | 520
[5 empty] | TE (ms) |  |  | 4.92
TI (ms)
[5 empty] | NSA |  |  | 1
[5 empty] | Scan duration | 0:01 |  | 0:01
[5 empty] | Peripheral gated (Y/N) | N |  | N
|  |  |  | Resolution | Matrix size: in-plane, freq dirn | 64 |  | 64
[5 empty] | Matrix size: in-plane, phase dirn | 64 |  | 64
[5 empty] | Total No. of slices (2D) or Matrix size: slice dirn (3D) | 36 |  | 36
[5 empty] | FOV frequency (cm) | 256 |  | 256
[5 empty] | FOV phase (%) | 100 |  | 100
[5 empty] | FOV slice (cm) | 13.2
[5 empty] | Voxel size: freq dirn (mm) | 4 |  | 4
[5 empty] | Voxel size: phase dirn (mm) | 4 |  | 4
[5 empty] | Voxel size: slice dirn (mm) | 4 |  | 4
[5 empty] | Slice gap (mm) | 0 |  | 0
|  |  |  | Coverage | 2D/3D |  |  | 2D
[5 empty] | In-plane phase enc. dirn. (AP,LR,SI) |  |  | R L
[5 empty] | Region of slice coverage | Whole Brain |  | WHOLE BRAIN
[5 empty] | Slice orientation (Sag/Ax/Cor/Obl) |  |  | AXIAL
[5 empty] | Slice order (Inter/Seq) |  |  | I TO S
Slice acquisition direction
K-space ordering
Concatenations
Oversampling Phase (%)
Oversampling Slice dirn 3D (%)
Slice overlap (mm)
|  |  |  | Sequence specific | ETL/ E Spacing (us)
TSE corrections
[5 empty] | Fat Sat (CHESS or Equiv) on (Y/N) |  |  | N
|  |  |  | Other acquisition | Partial Fourier (Y/N) |  |  | N
Partial Fourier method (NEX/Echo)
Fractional NEX
Fractional Echo
Flip angle (degrees)
rBW (Hz per pixel)
rBW (kHz)
|  |  |  | Parallel imaging | PI method |  |  | N
PI reduction factor
|  |  |  | Corrections | B0 mapping (Y/N)
B1 mapping: Transmit (Y/N)
B1 mapping: Receive (Y/N)
Grad Non-lin correction (Y/N)
Grad Non-lin correction Method
|  |  |  | Pre-processing | Zero filling (Y/N)